AON CTD Calibration HLY1003 HLY1103 and HLY1203

Manufacturer Calibrations

Calibrations for CTD sensors were determined for each CTD sensor (temperature, conductivity, pressure, etc.) by the manufacturer before the cruise. When appropriate, sensor calibration files and firmware were updated accordingly.

SeaBird Processing

As per manufacturer recommendations, CTD data were processed using SeaBird data processing software (ver. 7.22.0). The raw CTD data were converted from HEX to ASCII, lag corrected, edited for large spikes, smoothed according to sensor, and pressure averaged into 2 dB bins for final data quality control and analysis. The following table summarizes the processing routines used together with SeaBird-recommended parameters for the sensor configuration used.

| SeaBird Module | Description (SeaBird, Version 7.22.0) |
| --- | --- |
| DATCNV | Convert the raw data to pressure, temperature, conductivity, and dissolved oxygen current |
| ALIGNCTD | Advance conductivity approximately 0.073 seconds relative to pressure |
| WILDEDIT | Checks for and marks 'wild' data points: first pass 2.0 standard deviations; second pass 20 standard deviations |
| CELLTM | Conductivity cell thermal mass correction alpha = 0.03 and 1/beta = 7.0 |
| FILTER | Low pass filter pressure and depth with a time constant of 0.15 seconds to increase pressure resolution for LOOPEDIT |
| LOOPEDIT | Mark scans where the CTD is moving less than the minimum velocity (0.1 m/s) or traveling backwards due to ship roll |
| DERIVE oxy | Compute oxygen from oxygen current (filtered), temperature, and pressure |
| BINAVG | Average data into the 1 dbar pressure bins |
| DERIVE sal | Compute salinity |
| STRIP | Extract columns of data from .cnv files |
| SPLIT | Split .cnv file into upcast and downcast files |

Salinity Water Samples and Post Calibration

Water samples were collected from most stations occupied during this cruise. Samples were analyzed for concentrations and used to post-calibrate the CTD sensors.

WHOI post-processing fitting procedures are modeled after methods used in Millard and Yang, 1993. CTD conductivity and water sample salinity differences were characterized as a function of pressure and time. One fit was created by grouping together data from all CTD stations occupied. The group was fit for a slope and bias adjustment using only water sample data that was within a defined physical range of CTD values. The slope term is a polynomial function of the station number based upon chronological station collection order. A linear pressure term (modified beta) was applied to conductivity slopes using a least-squares minimization of CTD and bottle conductivity differences.

The function minimized was:

BC 	- bottle conductivity [mS/cm]

CC	- pre-cruise calibrated CTD conductivity [mS/cm]

CP	- CTD pressure [dbar]

m	- conductivity slope

b	- conductivity bias [mS/cm]

β	- linear pressure term [mS/cm/dbar]

The final conductivity, FC [mS/cm] is: